## Slide 1: Measurement Based IBIS Models
Tom Dagostino
VP Modeling
Teraspeed Consulting Group LLC
Device Modeling Division

## Slide 2: Agenda
What to expect
What does it take to make measurement based models
Process
Question

## Slide 3: What to Expect
Any component can be modeled
  No NDA required
  No SPICE deck required
  No compatibility issues
All you need is a sample and a datasheet
Notes: It is easy to get samples of the component that needs to be modeled unless the component is not in production yet. In most cases you do not need to sign a NDA to get samples or a datasheet.
With SPICE you frequently need to get a NDA signed which increases the time and cost of making a model. Lawyers do not do their work for free.
Many semiconductor vendors have proprietary SPICE simulators. Modeling a component which has a proprietary description means you will need a copy of the proprietary SPICE simulator.

## Slide 4: Accuracy
DC better than 1%
AC Bandwidth >10 GHz
AC accuracy limited by:
  Probing technique
  Fixturing
Notes: Its easy to measure the DC characteristics to the level of accuracy required. Getting good AC bandwidth requires good fixturing, good probing and high bandwidth samplers.

## Slide 5: IBIS Features Supported
Anything you can do with SPICE
  Have not tried
    Dynamic clamps
    Bus hold
    TT
Doing a fully coupled package model would be costly
Full min/typ/max if process corner components are available
Notes: Most IBIS features are described with either IV and VT tables or by taking the data right off the spec sheet. Dynamic clamps and Bus Hold require special setups and measurements. TT is easy to measure. We have not seen any demand for these features.

## Slide 6: Required Equipment
Wide bandwidth scope and specialized probes
  > 10 GHz
  50 Ohms load terminated to an arbitrary voltage
  TDR capability very useful
Notes: The bandwidth of the scope should be at least 5 times the bandwidth of the signals you model. Most parts these days have 200 to 500 psec transitions. This implies a scope bandwidth of at least 8.75 GHz.

## Slide 7
Notes: This is one of Teraspeed’s test setups. There are 5 power supplies, one scope with 20 GHz sampling heads and a 24 channel word generator. Also pictured is a hot/cold plate to vary the case temperature from 0 to 100C. There are a lot of other tools required such as probes, attenuators, etc.

## Slide 8: Stimulus
High speed word generator
  24 channels at 100 MHz
  Variable output levels
    ECL
    PECL
    LVDS
    CML

## Slide 9: Power Supplies
One supply to do IV sweeps
  At least +/- 7 volts
  At least +/- 1 Ampere
At least 4 additional supplies to power chip
One supply for termination voltage
Notes: Most outputs will source of sink between 100 and 500 ma. We have seen parts that can sink over 1 Amp. The supplies need to be able to both source or sink current to make these measurements.
Many of today’s chips have 3 or more supply connections. A core supple, a I/O supply or two and maybe some bias supplies.

## Slide 10: Misc. Equipment
Hot/cold plate 0 to 100 C
DMM
Device programmers
Signal generator
Temperature probe
Etc.

## Slide 11: SW Tools
SW to run the T&M equipment
  Takes all measurements and saves data
  Allows inspection, QC, measurements, etc.
SW to reduce the raw data to IBIS tables
  Extract clamps where needed
  Extrapolate endpoints
SW to write the IBIS model
  Automate the conversion of component data to IBIS files
Notes: Like most things the more automation you can use the more productive you can be. Also properly designed tools will prevent errors form occurring. Many of the problems seen in IBIS models are just syntax issues. Automated tools will prevent those kinds of mistakes.

## Slide 12: Measurement Process
VT data
  50 Ohms to ground and acquire rise and fall
  50 Ohms to Vcc and acquire rise and fall
  Preserve timing relationships between waveforms
Notes: Keeping the time correlation between waveforms is very important. The scope must be triggered off some external event, not the waveforms of interest.

## Slide 13
Notes: Here is a complete set of VT curves for a part measured over process, temperature and voltage. The upper set of waveforms are with the 50 Ohm termination tied to Vcc and the lower set for the output terminated to ground. It is easy to see the change in prop delay in the component from the max to typ to minimum set of conditions.

## Slide 14: Measurement Process
IV Data
  Place output in high state
    Sweep from 0 to Vcc + 1.2 Volts
  Place output in low state
    Sweep from –1.2V to Vcc
  If appropriate get input or tri-state
    Sweep from –1.2 to Vcc + 1.2 Volts

## Slide 15
Notes: These are the IV characteristics of the same component over process, temperature and voltage. The upper window shows the pull up characteristics, the middle the pulldown and the lower the input.

## Slide 16: Collect Component Data
Pin data
  Number
  Signal
  Model
  Package parasitics
  Electrical specification
  Timing test load
  Diff pin pairs
  Etc.

## Slide 17
Notes: This shows some of the screens in our database. All pin and specification data is stored in the database.

## Slide 18: Build Model
Build model
Check with IBIS Golden Parser
Load and check in a simulator
Notes: Too may model makers neglect step two in the process. There are too many models that will not pass the IBIS parser.

## Slide 19: Skills Needed
Thorough knowledge of logic
High frequency measurement techniques
More than passing knowledge of the IBIS spec.

## Slide 20: Question
Would a model with temperature and voltage variation be more useful than a typical only model?
Notes: Would a model that has voltage and temperature variation only, no process information give the SI engineer more useful information than a typical only model.